## according to ID
APPENDIX.3 : Test results reported by participating laboratories (Panel 19)
I.D. NO. | Assay | LOT NO. | HIV19/1 | HIV19/2 | HIV19/3 | HIV19/4 | HIV19/5 | HIV19/6 | HIV19/7 | HIV19/8
|  |  | NR | NR | R | NR | NR | R | NR | NR
1 | ABBOTT ARCHITECT | 0217BE00G | 0.12 | 0.11 | 786.34 | 0.16 | 0.09 | 800.27 | 0.13 | 0.13
2 | MERILISA HIV 1-2 Gen 3 | MI201938 | 0.095 | 0.063 | 10.059 | 0.059 | 0.086 | 10.095 | 0.059 | 0.063
4 | MERILISA HIV 1-2 Gen 3 | MI 101825 | 0.146 | 0.176 | 9.708 | 0.15 | 0.206 | 10.459 | 0.107 | 0.163
9 | MERILISA HIV 1-2 Gen 3 | MI 0 11928 | 0.124 | 0.115 | 9.597 | 0.097 | 0.102 | 9.235 | 0.133 | 0.164
9 | GENSCREEN ULTRA HIV Ag-Ab | 8H0123 | HIV19/3: 9.097; HIV19/6: 9.479
10 | MERILISA HIV 1-2 Gen 3 | MI1011928 | 0.046 | 0.046 | 17.545 | 0.064 | 0.06 | 17.358 | 0.064 | 0.069
11 | ABBOTT ARCHITECT | 03257BE00 | 0.19 | 0.12 | 925.29 | 0.18 | 0.12 | 901.31 | 0.1 | 0.19
11 | ORTHO CMIA VITROS | 390 | 0.27 | 0.15 | 168 | 0.12 | 0.15 | 160 | 0.14 | 0.13
14 | GENSCREEN ULTRA HIV Ag-Ab | 8K0770 | 0.325 | 0.358 | 9.771 | 0.293 | 0.312 | 9.771 | 0.315 | 0.351
15 | LABSYSTEM DIAG. TRIVITRON | HIVFE007 | 0.1172 | 0.0763 | 10.055 | 0.4127 | 0.0872 | 10.2727 | 0.0736 | 0.089
15 | HIV TRIDOT | HTD091829 | HIV19/3: REATIVE; HIV19/6: REATIVE
16 | MICROLISA J.MITRA Ag.& Ab. | EIA 020519 | 0.553 | 0.399 | 12.99 | 0.667 | 0.264 | 12.78 | 0.256 | 0.276
16 | QUALISA TULIP | 54393 | 0.033 | 0.004 | 13.112 | 0.004 | 0.0625 | 14.99 | 0.0384 | 0.067
18 | ROCHE COBAS E411 | 391001 | 0.327 | 0.262 | 772.11 | 0.27 | 0.27 | 760.33 | 0.281 | 0.293
18 | ROCHE COBAS E411 | 400139 | 0.295 | 0.236 | 694.9 | 0.243 | 0.243 | 684.3 | 0.253 | 0.264
26 | MERILISA HIV 1-2 Gen 3 | MI011928 | 0.085 | 0.098 | 7.156 | 0.085 | 0.13 | 7.941 | 0.094 | 0.098
30 | MERILISA HIV 1-2 Gen 3 | MI011928 | 0.09 | 0.157 | 8.649 | 0.209 | 0.177 | 7.952 | 0.169 | 0.165
32 | ABBOTT ARCHITECT | 03257BE00 | 0.179 | 0.111 | 802.211 | 0.107 | 0.081 | 778.352 | 0.144 | 0.121
32 | ABBOTT ARCHITECT | 03257BE00 | HIV19/3: 795.952; HIV19/6: 779.992
APPENDIX.3 : Test results reported by participating laboratories (Panel 19)
I.D. NO. | Assay | LOT NO. | HIV19/1 | HIV19/2 | HIV19/3 | HIV19/4 | HIV19/5 | HIV19/6 | HIV19/7 | HIV19/8
|  |  | NR | NR | R | NR | NR | R | NR | NR
33 | MERILISA HIV 1-2 Gen 3 | MI101864 | 0.113 | 0.097 | 9.105 | 0.097 | 0.168 | 9.164 | 0.118 | 0.13
34 | MERILISA HIV 1-2 Gen 3 | MI011938 | 0.26 | 0.24 | 12.77 | 0.17 | 0.37 | 11.98 | 0.3 | 0.19
36 | QUALISA TULIP | 49188 | 0.11 | 0.16 | 10.87 | 0.2 | 0.23 | 10.33 | 0.05 | 0.02
37 | MERILISA HIV Ag.& Ab. 4th GEN | MI011935 | 0.06 | 0.07 | 12.4 | 0.08 | 0.13 | 13.91 | 0.07 | 0.07
40 | MICROLISA J.MITRA | EIV010319 | 0.036 | 0.0053 | 4.8794 | 0.0197 | 0.02 | 4 | 0.0071 | 0.0089
40 | SIGNAL ARKRAY DIAG. | 4000020643 | NAGATIVE | NAGATIVE | POSITIVE | NAGATIVE | NAGATIVE | POSITIVE | NAGATIVE | NAGATIVE
41 | MERILISA HIV 1-2 Gen 3 | MI101850 | 0.099 | 0.163 | 8.801 | 0.129 | 0.185 | 9.672 | 0.099 | 0.051
41 | MERILISA HIV 1-2 Gen 3 | MI101850 | 0.168 | 0.141 | 9.566 | 0.11 | 0.159 | 9.898 | 0.119 | 0.101
43 | ROCHE COBAS E411 | 391001 | 0.297 | 0.325 | 776.2 | 0.298 | 0.295 | 711.9 | 0.283 | 0.669
45 | MERILISA HIV 1-2 Gen 3 | MI101827 | 0.08 | 0.09 | 13.31 | 0.12 | 0.11 | 13.24 | 0.13 | 0.09
50 | MERILISA HIV 1-2 GEN 3 | MI011928 | 0.105 | 0.155 | 13.61 | 0.037 | 0.046 | 12.59 | 0.088 | 0.214
52 | MERILISA HIV 1-2 Gen 3 | MI011928 | 0.132 | 0.0334 | 3.93 | 0.0271 | 0.0186 | 3.83 | 0.0728 | 0.29
52 | DIAGNOSTIC ENTERPRISES | HTD061922 | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE
55 | GENSCREEN ULTRA HIV Ag-Ab | 8K0770 | 0.356 | 0.372 | 28.662 | 0.347 | 0.347 | 10.98 | 0.318 | 0.375
55 | DIAGNOSTIC ENTERPRISES | HTD121835 | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE
56 | QUALISA TULIP | 49187 | 0.141 | 0.114 | 13.882 | 0.136 | 0.123 | 13.923 | 0.136 | 0.105
57 | MICROLSA J.MITRA Ag.& Ab. | EIA 020519 | 0.019 | 0.048 | 2.729 | 0.049 | 0.038 | 2.916 | 0.032 | 0.043
57 | ALERE HIV1/2 | 92758K100A | N.R | N.R | REACTIVE | N.R | N.R | REACTIVE | N.R | N.R
APPENDIX.3 : Test results reported by participating laboratories (Panel 19)
I.D. NO. | Assay | LOT NO. | HIV19/1 | HIV19/2 | HIV19/3 | HIV19/4 | HIV19/5 | HIV19/6 | HIV19/7 | HIV19/8
|  |  | NR | NR | R | NR | NR | R | NR | NR
58 | MERILISA HIV Ag.& Ab. 4th GEN | MI051952 | 0.045 | 0.068 | 6.137 | 0.114 | 0.064 | 4.646 | 0.022 | 0.059
58 | HIV TRIDOT | HTD091828 | HIV19/3: REACTIVE; HIV19/6: REACTIVE
60 | QUALISA TULIP | 49189 | 0.141 | 0.22 | 13.137 | 0.163 | 0.15 | 12.115 | 0.009 | 0.022
60 | HIV TRIDOT | HTD041911 | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE
61 | MERILISA HIV Ag.& Ab. 4th GEN | MI051952 | 0.063 | 0.058 | 9.603 | 0.054 | 0.049 | 9.337 | 0.045 | 0.049
61 | TRIDOT TULIP DIAG. | HT0051915 | NEGATIVE | NEGATIVE | POSITIVE | NEGATIVE | NEGATIVE | POSITIVE | NEGATIVE | NEGATIVE
64 | MERILISA HIV 1-2 Gen 3 | MI101826 | 0.2132 | 0.2182 | 16.662 | 0.2083 | 0.238 | 15.912 | 0.2331 | 0.2132
64 | RETROQUIC TULIP DIAG. | 541241 | N.R | N.R | REACTIVE | N.R | N.R | REACTIVE | N.R | N.R
67 | ABBOTT ARCHITECT | 02175BE00 | 0.13 | 0.1 | 766.44 | 0.13 | 0.11 | 782.45 | 0.09 | 0.11
68 | GENSCREEN ULTRA HIV Ag-Ab | 9A0772 | 0.2554 | 0.2595 | 14.18 | 0.2108 | 0.2554 | 14.18 | 0.2554 | 0.223
68 | HIV TRIDOT | HTD111833 | HIV19/3: REACTIVE; HIV19/6: REACTIVE
69 | QUALISA TULIP | 49186 | 0.101 | 0.106 | 20.85 | 0.131 | 0.035 | 20.6 | 0.23 | 0.091
69 | HIV TRIDOT | HTD121835 | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE
70 | ABBOTT ARCHITECT | 02175BE00 | 0.14 | 0.13 | 822.14 | 0.16 | 0.1 | 833.81 | 0.11 | 0.11
73 | MERILISA HIV 1-2 Gen 3 | MI011927 | 0.005 | 0.005 | 12.7 | 0.03 | 0.005 | 13.02 | 0.005 | 0.003
73 | COMBAIDS ARKRAY | 4000021345 | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE
74 | ORTHO CLINACAL DIAG. | 390 | 0.11 | 0.09 | 188 | 0.1 | 0.05 | 151 | 0.08 | 0.1
74 | MICROLSA J.MITRA Ag.& Ab. | EIA040818 | 0.12 | 0.06 | 8.9 | 0.16 | 0.041 | 7.87 | 0.06 | 0.07
74 | DIAGNOSTIC ENTERPRISES | HTD121836 | NEGATIVE | NEGATIVE | RECTIVE | NEGATIVE | NEGATIVE | RECTIVE | NEGATIVE | NEGATIVE
75 | ORTHO CLINACAL DIAG. | 390 | 0.1 | 0.07 | 166 | 0.08 | 0.08 | 169 | 0.07 | 0.1
APPENDIX.3 : Test results reported by participating laboratories (Panel 19)
I.D. NO. | Assay | LOT NO. | HIV19/1 | HIV19/2 | HIV19/3 | HIV19/4 | HIV19/5 | HIV19/6 | HIV19/7 | HIV19/8
|  |  | NR | NR | R | NR | NR | R | NR | NR
78 | BENESPHERA AVANTOR | DBHIA0119A | 0.089 | 0.073 | 13.277 | 0.082 | 0.089 | 13.916 | 0.082 | 0.087
79 | MERILISA HIV 1-2 Gen 3 | MI021923 | 0.058 | 0.102 | 9.457 | 0.071 | 0.043 | 7.793 | 0.062 | 0.084
82 | MERILISA HIV 1-2 Gen 3 | MI101864 | 0.101 | 0.115 | 10.74 | 0.101 | 0.119 | 11.163 | 0.123 | 0.097
85 | QUALISA QULPRO TULIP | 49188 | 0.208 | 0.323 | 10.169 | 0.226 | 0.329 | 10.234 | 0.303 | 0.216
87 | QUALISA QULPRO TULIP | 49180 | 0.011 | 0.255 | 12.962 | 0.04 | 0.003 | 12.962 | 0.355 | 0.351
87 | QUALPRO DIAG. | 48426 | HIV19/3: REACTIVE; HIV19/6: REACTIVE
88 | ABBOTT ARCHITECT | 02175BE00 | 0.12 | 0.1 | 807.1 | 0.09 | 0.12 | 794.17 | 0.07 | 0.08
89 | MERILISA HIV 1-2 Gen 3 | MI041924 | 0.05213 | 0.04739 | 10.1706 | 0.03791 | 0.04739 | 10.8625 | 0.042654 | 0.0284
89 | MERSCREEN HIV 1-2 WB MERIL | MI081855 | HIV19/3: REACTIVE; HIV19/6: REACTIVE
90 | ORTHO VITROS ECI | 370 | 0.14 |  | 132 | 0.15 | 0.12 | 136 | 0.12 | 0.15
91 | ORTHO VITROS ECI | 360 | 0.11 | 0.08 | 162 | 0.08 | 0.13 | 179 | 0.14 | 0.16
93 | MERILISA HIV 1-2 Gen 3 | MI011928 | 0.155 | 0.189 | 13.411 | 0.193 | 0.487 | 12.924 | 0.638 | 0.218
93 | SD DAIGNOSTIC | 03DDD010A | NON-REACTIVE | NON-REACTIVE | REACTIVE | NON-REACTIVE | NON-REACTIVE | REACTIVE | NON-REACTIVE | NON-REACTIVE
95 | ROCHE COBAS E411 | 391001 | 0.318 | 0.294 | 567.7 | 0.311 | 0.302 | 636.5 | 0.29 | 0.304
95 | ROCHE COBAS E411 | 391001 | HIV19/3: 559.1; HIV19/6: 625.4
96 | MERILISA HIV 1-2 Gen 3 | MI011937 | 0.043 | 0.038 | 15.251 | 0.038 | 0.043 | 14.754 | 0.014 | 0.038
96 | HIV TRIDOT | HTD031907 | HIV19/3: REACTIVE; HIV19/6: REACTIVE
APPENDIX.3 : Test results reported by participating laboratories (Panel 19)
I.D. NO. | Assay | LOT NO. | HIV19/1 | HIV19/2 | HIV19/3 | HIV19/4 | HIV19/5 | HIV19/6 | HIV19/7 | HIV19/8
|  |  | NR | NR | R | NR | NR | R | NR | NR
97 | MERILISA HIV 1-2 Gen 3 | MI041924 | 0.132 | 0.123 | 6.229 | 0.115 | 0.123 | 7.379 | 0.119 | 0.123
97 | HIV TRDOT+ Ag. 4th GEN J.MITRA | HTA010119 | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE
98 | ORTHO CLINACAL DIAG. | 370 | 0.2 | 0.18 | 137 | 0.17 | 0.13 | 83.1 | 0.18 | 0.19
98 | HIV TRIDOT | HTD030319 | HIV19/3: REACTIVE; HIV19/6: REACTIVE
98 | ELFA HIV DUO ULTRA BIOMEREUX | 1007062230 | HIV19/3: 24.12; HIV19/6: 12.77
101 | ORTHO CLINACAL DIAG. | 390 | 0.25 | 0.24 | 173 | 0.23 | 0.19 | 170 | 0.21 | 0.23
101 | ABBOTT ARCHITECT |  | HIV19/3: 852.49; HIV19/6: 829.33
102 | ERBALISA TRANSASIA | TIG31806 | 0.281 | 0.187 | 23.562 | 0.25 | 0.21 | 20.835 | 0.14 | 0.281
102 | MERILISA HIV 1-2 Gen 3 | MI101850 | 0.017 | 0.017 | 11.674 | 0.022 | 0.035 | 11.857 | 0.013 | 0.075
103 | BENESPHERA AVANTOR | DHJAI0319 | 0.038 | 0.033 | 12.339 | 0.009 | 0.095 | 11.923 | 0.019 | 0.033
104 | ORTHO CLINACAL DIAG. | 370 | 0.08 | 0.07 | 152 | 0.06 | 0.06 | 149 | 0.06 | 0.08
110 | ORTHO CLINACAL DIAG. | 370 | 0.18 | 0.08 | 146 | 0.1 | 0.08 | 136 | 0.08 | 0.12
112 | GENSCREEN ULTRA HIV Ag-Ab | 8H0123 | 0.1515 | 0.1645 | 12.645 | 0.1558 | 0.1688 | 12.5541 | 0.1602 | 0.1429
113 | ORTTHO CLINICAL DIAG. | 360 | 0.16 | 0.1 | 171 | 0.13 | 0.14 | 173 | 0.17 | 0.25
115 | ELECTRA TULIP | 410-016 | 0.052 | 0.058 | 8.983 | 0.076 | 0.072 | 8.989 | 0.052 | 0.099
115 | BHATBIOTECH AIDSCAN | 37190407 | NEGATIVE | NEGATIVE | POSITIVE | NEGATIVE | NEGATIVE | POSITIVE | NEGATIVE | NEGATIVE
117 | OSCAR MEDICARE | HIV4EL-010 | 0.2124 | 0.1983 | 11.6465 | 0.2232 | 0.1871 | 12.8112 | 0.1626 | 0.1654
117 | DIAGNOSTIC ENTERPRISES | HTD041911 | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE
APPENDIX.3 : Test results reported by participating laboratories (Panel 19)
I.D. NO. | Assay | LOT NO. | HIV19/1 | HIV19/2 | HIV19/3 | HIV19/4 | HIV19/5 | HIV19/6 | HIV19/7 | HIV19/8
|  |  | NR | NR | R | NR | NR | R | NR | NR
120 | ORTHO CLINACAL DIAG. | 390 | 0.11 | 0.08 | 161 | 0.09 | 0.09 | 175 | 0.11 | 0.13
121 | ORTHO VITROS ECI | 360 | 0.15 | 0.11 | 179 | 0.11 | 0.13 | 176 | 0.13 | 0.14
122 | MICROLSA J.MITRA Ag.& Ab. | EIA 020519 | 0.328 | 0.306 | 9.91 | 0.164 | 0.213 | 11.293 | 0.24 | 0.182
123 | ORTHO CLINACAL DIAG. | 360 | 0.11 | 0.13 | 200 | 0.1 | 0.08 | 201 | 0.08 | 0.1
124 | CHEMILUMINESCENCE QUALPRO DIAG. | 410-016 | 0.068 | 0.082 | 8.192 | 0.092 | 0.075 | 8.48 | 0.075 | 0.08
125 | BHAT BIOTECH HIV I / II | 37190103 | 0.0613 | 0.04192 | 5.4427 | 0.0148 | 0.0155 | 5.7205 | 0.0196 | 0.06637
125 | BHAT BIOTECH AIDSCAN | 79180902 | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE
126 | GENSCREEN ULTRA HIV Ag-Ab |  | 0.229 | 0.212 | 11.98 | 0.2089 | 0.212 | 11.551 | 0.188 | 0.274
126 | COMBAIDS ARKRAY | 4000021345 | N.R | N.R | REACTIVE | N.R | N.R | REACTIVE | N.R | N.R
127 | CHEMILUMINESCENCE QUALPRO DIAG. | 410-014 | 0.068 | 0.063 | 12.907 | 0.081 | 0.084 | 12.97 | 0.09 | 0.064
128 | ORTHO VITROS ECI | 360 | 0.16 | 0.08 | 194 | 0.11 | 0.12 | 198 | 0.18 | 0.11
129 | QUALISA TULIP | 49191 | 0.07 | 0.07 | 8.744 | 0.07 | 0.065 | 8.107 | 0.065 | 0.07

## ACCORDING TO ASSAY
I.D. NO. | Assay | LOT NO. | HIV19/1 | HIV19/2 | HIV19/3 | HIV19/4 | HIV19/5 | HIV19/6 | HIV19/7 | HIV19/8
1 | ABBOTT ARCHITECT | 0217BE00G | 0.12 | 0.11 | 786.34 | 0.16 | 0.09 | 800.27 | 0.13 | 0.13
11 | ABBOTT ARCHITECT | 03257BE00 | 0.19 | 0.12 | 925.29 | 0.18 | 0.12 | 901.31 | 0.1 | 0.19
32 | ABBOTT ARCHITECT | 03257BE00 | 0.179 | 0.111 | 802.211 | 0.107 | 0.081 | 778.352 | 0.144 | 0.121
32 | ABBOTT ARCHITECT | 03257BE00 | HIV19/3: 795.952; HIV19/6: 779.992
67 | ABBOTT ARCHITECT | 02175BE00 | 0.13 | 0.1 | 766.44 | 0.13 | 0.11 | 782.45 | 0.09 | 0.11
70 | ABBOTT ARCHITECT | 02175BE00 | 0.14 | 0.13 | 822.14 | 0.16 | 0.1 | 833.81 | 0.11 | 0.11
88 | ABBOTT ARCHITECT | 02175BE00 | 0.12 | 0.1 | 807.1 | 0.09 | 0.12 | 794.17 | 0.07 | 0.08
101 | ABBOTT ARCHITECT |  | HIV19/3: 852.49; HIV19/6: 829.33
57 | ALERE HIV1/2 | 92758K100A | bar: HIV19/1 to HIV19/8
78 | BENESPHERA AVANTOR | DBHIA0119A | 0.089 | 0.073 | 13.277 | 0.082 | 0.089 | 13.916 | 0.082 | 0.087
103 | BENESPHERA AVANTOR | DHJAI0319 | 0.038 | 0.033 | 12.339 | 0.009 | 0.095 | 11.923 | 0.019 | 0.033
125 | BHAT BIOTECH AIDSCAN | 79180902 | bar: HIV19/1 to HIV19/8
115 | BHAT BIOTECH AIDSCAN | 37190407 | bar: HIV19/1 to HIV19/8
125 | BHAT BIOTECH HIV I / II | 37190103 | 0.0613 | 0.04192 | 5.4427 | 0.0148 | 0.0155 | 5.7205 | 0.0196 | 0.06637
124 | CHEMILUMINESCENCE QUALPRO DIAG. | 410-016 | 0.068 | 0.082 | 8.192 | 0.092 | 0.075 | 8.48 | 0.075 | 0.08
127 | CHEMILUMINESCENCE QUALPRO DIAG. | 410-014 | 0.068 | 0.063 | 12.907 | 0.081 | 0.084 | 12.97 | 0.09 | 0.064
73 | COMBAIDS ARKRAY | 4000021345 | bar: HIV19/1 to HIV19/8
126 | COMBAIDS ARKRAY | 4000021345 | bar: HIV19/1 to HIV19/8
52 | DIAGNOSTIC ENTERPRISES | HTD061922 | bar: HIV19/1 to HIV19/8
55 | DIAGNOSTIC ENTERPRISES | HTD121835 | bar: HIV19/1 to HIV19/8
74 | DIAGNOSTIC ENTERPRISES | HTD121836 | bar: HIV19/1 to HIV19/8
117 | DIAGNOSTIC ENTERPRISES | HTD041911 | bar: HIV19/1 to HIV19/8
115 | ELECTRA TULIP | 410-016 | 0.052 | 0.058 | 8.983 | 0.076 | 0.072 | 8.989 | 0.052 | 0.099
98 | ELFA HIV DUO ULTRA BIOMEREUX | 1007062230 | HIV19/3: 24.12; HIV19/6: 12.77
102 | ERBALISA TRANSASIA | TIG31806 | 0.281 | 0.187 | 23.562 | 0.25 | 0.21 | 20.835 | 0.14 | 0.281
9 | GENSCREEN ULTRA HIV Ag-Ab | 8H0123 | HIV19/3: 9.097; HIV19/6: 9.479
14 | GENSCREEN ULTRA HIV Ag-Ab | 8K0770 | 0.325 | 0.358 | 9.771 | 0.293 | 0.312 | 9.771 | 0.315 | 0.351
55 | GENSCREEN ULTRA HIV Ag-Ab | 8K0770 | 0.356 | 0.372 | 28.662 | 0.347 | 0.347 | 10.98 | 0.318 | 0.375
68 | GENSCREEN ULTRA HIV Ag-Ab | 9A0772 | 0.2554 | 0.2595 | 14.18 | 0.2108 | 0.2554 | 14.18 | 0.2554 | 0.223
112 | GENSCREEN ULTRA HIV Ag-Ab | 8H0123 | 0.1515 | 0.1645 | 12.645 | 0.1558 | 0.1688 | 12.5541 | 0.1602 | 0.1429
126 | GENSCREEN ULTRA HIV Ag-Ab |  | 0.229 | 0.212 | 11.98 | 0.2089 | 0.212 | 11.551 | 0.188 | 0.274
97 | HIV TRDOT+ Ag. 4th GEN J.MITRA | HTA010119 | bar: HIV19/1 to HIV19/8
15 | HIV TRIDOT | HTD091829 | bar: HIV19/3; HIV19/6
58 | HIV TRIDOT | HTD091828 | bar: HIV19/3; HIV19/6
60 | HIV TRIDOT | HTD041911 | bar: HIV19/1 to HIV19/8
68 | HIV TRIDOT | HTD111833 | bar: HIV19/3; HIV19/6
69 | HIV TRIDOT | HTD121835 | bar: HIV19/1 to HIV19/8
96 | HIV TRIDOT | HTD031907 | bar: HIV19/3; HIV19/6
98 | HIV TRIDOT | HTD030319 | bar: HIV19/3; HIV19/6
15 | LABSYSTEM DIAG. TRIVITRON | HIVFE007 | 0.1172 | 0.0763 | 10.055 | 0.4127 | 0.0872 | 10.2727 | 0.0736 | 0.089
2 | MERILISA HIV 1-2 Gen 3 | MI201938 | 0.095 | 0.063 | 10.059 | 0.059 | 0.086 | 10.095 | 0.059 | 0.063
4 | MERILISA HIV 1-2 Gen 3 | MI 101825 | 0.146 | 0.176 | 9.708 | 0.15 | 0.206 | 10.459 | 0.107 | 0.163
9 | MERILISA HIV 1-2 Gen 3 | MI 0 11928 | 0.124 | 0.115 | 9.597 | 0.097 | 0.102 | 9.235 | 0.133 | 0.164
10 | MERILISA HIV 1-2 Gen 3 | MI1011928 | 0.046 | 0.046 | 17.545 | 0.064 | 0.06 | 17.358 | 0.064 | 0.069
26 | MERILISA HIV 1-2 Gen 3 | MI011928 | 0.085 | 0.098 | 7.156 | 0.085 | 0.13 | 7.941 | 0.094 | 0.098
30 | MERILISA HIV 1-2 Gen 3 | MI011928 | 0.09 | 0.157 | 8.649 | 0.209 | 0.177 | 7.952 | 0.169 | 0.165
33 | MERILISA HIV 1-2 Gen 3 | MI101864 | 0.113 | 0.097 | 9.105 | 0.097 | 0.168 | 9.164 | 0.118 | 0.13
34 | MERILISA HIV 1-2 Gen 3 | MI011938 | 0.26 | 0.24 | 12.77 | 0.17 | 0.37 | 11.98 | 0.3 | 0.19
41 | MERILISA HIV 1-2 Gen 3 | MI101850 | 0.099 | 0.163 | 8.801 | 0.129 | 0.185 | 9.672 | 0.099 | 0.051
41 | MERILISA HIV 1-2 Gen 3 | MI101850 | 0.168 | 0.141 | 9.566 | 0.11 | 0.159 | 9.898 | 0.119 | 0.101
45 | MERILISA HIV 1-2 Gen 3 | MI101827 | 0.08 | 0.09 | 13.31 | 0.12 | 0.11 | 13.24 | 0.13 | 0.09
52 | MERILISA HIV 1-2 Gen 3 | MI011928 | 0.132 | 0.0334 | 3.93 | 0.0271 | 0.0186 | 3.83 | 0.0728 | 0.29
64 | MERILISA HIV 1-2 Gen 3 | MI101826 | 0.2132 | 0.2182 | 16.662 | 0.2083 | 0.238 | 15.912 | 0.2331 | 0.2132
73 | MERILISA HIV 1-2 Gen 3 | MI011927 | 0.005 | 0.005 | 12.7 | 0.03 | 0.005 | 13.02 | 0.005 | 0.003
79 | MERILISA HIV 1-2 Gen 3 | MI021923 | 0.058 | 0.102 | 9.457 | 0.071 | 0.043 | 7.793 | 0.062 | 0.084
82 | MERILISA HIV 1-2 Gen 3 | MI101864 | 0.101 | 0.115 | 10.74 | 0.101 | 0.119 | 11.163 | 0.123 | 0.097
89 | MERILISA HIV 1-2 Gen 3 | MI041924 | 0.05213 | 0.04739 | 10.1706 | 0.03791 | 0.04739 | 10.8625 | 0.042654 | 0.0284
93 | MERILISA HIV 1-2 Gen 3 | MI011928 | 0.155 | 0.189 | 13.411 | 0.193 | 0.487 | 12.924 | 0.638 | 0.218
96 | MERILISA HIV 1-2 Gen 3 | MI011937 | 0.043 | 0.038 | 15.251 | 0.038 | 0.043 | 14.754 | 0.014 | 0.038
97 | MERILISA HIV 1-2 Gen 3 | MI041924 | 0.132 | 0.123 | 6.229 | 0.115 | 0.123 | 7.379 | 0.119 | 0.123
102 | MERILISA HIV 1-2 Gen 3 | MI101850 | 0.017 | 0.017 | 11.674 | 0.022 | 0.035 | 11.857 | 0.013 | 0.075
50 | MERILISA HIV 1-2 GEN 3 | MI011928 | 0.105 | 0.155 | 13.61 | 0.037 | 0.046 | 12.59 | 0.088 | 0.214
37 | MERILISA HIV Ag.& Ab. 4th GEN | MI011935 | 0.06 | 0.07 | 12.4 | 0.08 | 0.13 | 13.91 | 0.07 | 0.07
58 | MERILISA HIV Ag.& Ab. 4th GEN | MI051952 | 0.045 | 0.068 | 6.137 | 0.114 | 0.064 | 4.646 | 0.022 | 0.059
61 | MERILISA HIV Ag.& Ab. 4th GEN | MI051952 | 0.063 | 0.058 | 9.603 | 0.054 | 0.049 | 9.337 | 0.045 | 0.049
89 | MERSCREEN HIV 1-2 WB MERIL | MI081855 | bar: HIV19/3; HIV19/6
40 | MICROLSA J.MITRA | EIV010319 | 0.036 | 0.0053 | 4.8794 | 0.0197 | 0.02 | 4 | 0.0071 | 0.0089
16 | MICROLSA J.MITRA Ag.& Ab. | EIA 020519 | 0.553 | 0.399 | 12.99 | 0.667 | 0.264 | 12.78 | 0.256 | 0.276
57 | MICROLSA J.MITRA Ag.& Ab. | EIA 020519 | 0.019 | 0.048 | 2.729 | 0.049 | 0.038 | 2.916 | 0.032 | 0.043
74 | MICROLSA J.MITRA Ag.& Ab. | EIA040818 | 0.12 | 0.06 | 8.9 | 0.16 | 0.041 | 7.87 | 0.06 | 0.07
122 | MICROLSA J.MITRA Ag.& Ab. | EIA 020519 | 0.328 | 0.306 | 9.91 | 0.164 | 0.213 | 11.293 | 0.24 | 0.182
74 | ORTHO VITROS ECI | 390 | 0.11 | 0.09 | 188 | 0.1 | 0.05 | 151 | 0.08 | 0.1
75 | ORTHO VITROS ECI | 390 | 0.1 | 0.07 | 166 | 0.08 | 0.08 | 169 | 0.07 | 0.1
98 | ORTHO VITROS ECI | 370 | 0.2 | 0.18 | 137 | 0.17 | 0.13 | 83.1 | 0.18 | 0.19
101 | ORTHO VITROS ECI | 390 | 0.25 | 0.24 | 173 | 0.23 | 0.19 | 170 | 0.21 | 0.23
104 | ORTHO VITROS ECI | 370 | 0.08 | 0.07 | 152 | 0.06 | 0.06 | 149 | 0.06 | 0.08
110 | ORTHO VITROS ECI | 370 | 0.18 | 0.08 | 146 | 0.1 | 0.08 | 136 | 0.08 | 0.12
120 | ORTHO VITROS ECI | 390 | 0.11 | 0.08 | 161 | 0.09 | 0.09 | 175 | 0.11 | 0.13
123 | ORTHO VITROS ECI | 360 | 0.11 | 0.13 | 200 | 0.1 | 0.08 | 201 | 0.08 | 0.1
11 | ORTHO VITROS ECI | 390 | 0.27 | 0.15 | 168 | 0.12 | 0.15 | 160 | 0.14 | 0.13
90 | ORTHO VITROS ECI | 370 | 0.14 |  | 132 | 0.15 | 0.12 | 136 | 0.12 | 0.15
91 | ORTHO VITROS ECI | 360 | 0.11 | 0.08 | 162 | 0.08 | 0.13 | 179 | 0.14 | 0.16
121 | ORTHO VITROS ECI | 360 | 0.15 | 0.11 | 179 | 0.11 | 0.13 | 176 | 0.13 | 0.14
128 | ORTHO VITROS ECI | 360 | 0.16 | 0.08 | 194 | 0.11 | 0.12 | 198 | 0.18 | 0.11
113 | ORTHO VITROS ECI | 360 | 0.16 | 0.1 | 171 | 0.13 | 0.14 | 173 | 0.17 | 0.25
117 | OSCAR MEDICARE | HIV4EL-010 | 0.2124 | 0.1983 | 11.6465 | 0.2232 | 0.1871 | 12.8112 | 0.1626 | 0.1654
85 | QUALISA QULPRO TULIP | 49188 | 0.208 | 0.323 | 10.169 | 0.226 | 0.329 | 10.234 | 0.303 | 0.216
87 | QUALISA QULPRO TULIP | 49180 | 0.011 | 0.255 | 12.962 | 0.04 | 0.003 | 12.962 | 0.355 | 0.351
16 | QUALISA QULPRO TULIP | 54393 | 0.033 | 0.004 | 13.112 | 0.004 | 0.0625 | 14.99 | 0.0384 | 0.067
36 | QUALISA QULPRO TULIP | 49188 | 0.11 | 0.16 | 10.87 | 0.2 | 0.23 | 10.33 | 0.05 | 0.02
56 | QUALISA QULPRO TULIP | 49187 | 0.141 | 0.114 | 13.882 | 0.136 | 0.123 | 13.923 | 0.136 | 0.105
60 | QUALISA QULPRO TULIP | 49189 | 0.141 | 0.22 | 13.137 | 0.163 | 0.15 | 12.115 | 0.009 | 0.022
69 | QUALISA QULPRO TULIP | 49186 | 0.101 | 0.106 | 20.85 | 0.131 | 0.035 | 20.6 | 0.23 | 0.091
129 | QUALISA QULPRO TULIP | 49191 | 0.07 | 0.07 | 8.744 | 0.07 | 0.065 | 8.107 | 0.065 | 0.07
87 | QUALPRO DIAG. | 48426 | bar: HIV19/3; HIV19/6
64 | RETROQUIC TULIP DIAG. | 541241 | bar: HIV19/1 to HIV19/8
18 | ROCHE COBAS E411 | 391001 | 0.327 | 0.262 | 772.11 | 0.27 | 0.27 | 760.33 | 0.281 | 0.293
18 | ROCHE COBAS E411 | 400139 | 0.295 | 0.236 | 694.9 | 0.243 | 0.243 | 684.3 | 0.253 | 0.264
43 | ROCHE COBAS E411 | 391001 | 0.297 | 0.325 | 776.2 | 0.298 | 0.295 | 711.9 | 0.283 | 0.669
95 | ROCHE COBAS E411 | 391001 | 0.318 | 0.294 | 567.7 | 0.311 | 0.302 | 636.5 | 0.29 | 0.304
95 | ROCHE COBAS E411 | 391001 | HIV19/3: 559.1; HIV19/6: 625.4
93 | SD DAIGNOSTIC | 03DDD010A | bar: HIV19/1 to HIV19/8
40 | SIGNAL ARKRAY DIAG. | 4000020643 | bar: HIV19/1 to HIV19/8
61 | TRIDOT TULIP DIAG. | HT0051915 | bar: HIV19/1 to HIV19/8

## FOR STATISTICS
APPENDIX.3 : Test results reported by participating laboratories (Panel 19)
I.D. NO. | Assay | LOT NO. | HIV19/1 | HIV19/2 | HIV19/3 | HIV19/4 | HIV19/5 | HIV19/6 | HIV19/7 | HIV19/8
1 | ABBOTT ARCHITECT | 0217BE00G | 0.12 | 0.11 | 786.34 | 0.16 | 0.09 | 800.27 | 0.13 | 0.13
11 | ABBOTT ARCHITECT | 0.3257BE00 | 0.19 | 0.12 | 925.29 | 0.18 | 0.12 | 901.31 | 0.1 | 0.19
32 | ABBOTT ARCHITECT | 03257BE00 | 0.179 | 0.111 | 802.211 | 0.107 | 0.081 | 778.352 | 0.144 | 0.121
32 | ABBOTT ARCHITECT | 03257BE00 | HIV19/3: 795.952; HIV19/6: 779.992
67 | ABBOTT ARCHITECT | 02175BE00 | 0.13 | 0.1 | 766.44 | 0.13 | 0.11 | 782.45 | 0.09 | 0.11
70 | ABBOTT ARCHITECT | 02175BE00 | 0.14 | 0.13 | 822.14 | 0.16 | 0.1 | 833.81 | 0.11 | 0.11
88 | ABBOTT ARCHITECT | 02175BE00 | 0.12 | 0.1 | 807.1 | 0.09 | 0.12 | 794.17 | 0.07 | 0.08
101 | ABBOTT ARCHITECT |  | HIV19/3: 852.49; HIV19/6: 829.33
6 | 8 |  | 6 | 6 | 8 | 6 | 6 | 8 | 6 | 6
| Mean |  | 0.1465 | 0.11183 | 819.7454 | 0.13783 | 0.1035 | 812.4605 | 0.10733 | 0.1235
| Median |  | 0.135 | 0.1105 | 804.6555 | 0.145 | 0.105 | 797.22 | 0.105 | 0.1155
| S.Dev |  | 0.030553 | 0.011669 | 49.5818 | 0.034816 | 0.016047 | 41.8091 | 0.026882 | 0.036681
| C.V |  | 20.8554 | 10.4343 | 6.0484 | 25.2596 | 15.5042 | 5.146 | 25.0458 | 29.7013
| Upper Limit |  | 0.20761 | 0.13517 | 918.909 | 0.20747 | 0.13559 | 896.0786 | 0.1611 | 0.19686
| Lower Limit |  | 0.085394 | 0.088495 | 720.5818 | 0.068201 | 0.071406 | 728.8424 | 0.053568 | 0.050138
57 | ALERE HIV1/2 | 92758K100A | N.R | N.R | REACTIVE | N.R | N.R | REACTIVE | N.R | N.R
78 | BENESPHERA AVANTOR | DBHIA0119A | 0.089 | 0.073 | 13.277 | 0.082 | 0.089 | 13.916 | 0.082 | 0.087
103 | BENESPHERA AVANTOR | DHJAI0319 | 0.038 | 0.033 | 12.339 | 0.009 | 0.095 | 11.923 | 0.019 | 0.033
125 | BHAT BIOTECH AIDSCAN | 79180902 | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE
115 | BHAT BIOTECH AIDSCAN | 37190407 | NEGATIVE | NEGATIVE | POSITIVE | NEGATIVE | NEGATIVE | POSITIVE | NEGATIVE | NEGATIVE
125 | BHAT BIOTECH HIV I / II | 37190103 | 0.0613 | 0.04192 | 5.4427 | 0.0148 | 0.0155 | 5.7205 | 0.0196 | 0.06637
124 | CHEMILUMINESCENCE QUALPRO DIAG. | 410-016 | 0.068 | 0.082 | 8.192 | 0.092 | 0.075 | 8.48 | 0.075 | 0.08
127 | CHEMILUMINESCENCE QUALPRO DIAG. | 410-014 | 0.068 | 0.063 | 12.907 | 0.081 | 0.084 | 12.97 | 0.09 | 0.064
73 | COMBAIDS ARKRAY | 4000021345 | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE
126 | COMBAIDS ARKRAY | 4000021345 | N.R | N.R | REACTIVE | N.R | N.R | REACTIVE | N.R | N.R
115 | ELECTRA TULIP | 410-016 | 0.052 | 0.058 | 8.983 | 0.076 | 0.072 | 8.989 | 0.052 | 0.099
98 | ELFA HIV DUO ULTRA BIOMEREUX | 1007062230 | HIV19/3: 24.12; HIV19/6: 12.77
102 | ERBALISA TRANSASIA | TIG31806 | 0.281 | 0.187 | 23.562 | 0.25 | 0.21 | 20.835 | 0.14 | 0.281
9 | GENSCREEN ULTRA HIV Ag-Ab | 8H0123 | HIV19/3: 9.097; HIV19/6: 9.479
14 | GENSCREEN ULTRA HIV Ag-Ab | 8K0770 | 0.325 | 0.358 | 9.771 | 0.293 | 0.312 | 9.771 | 0.315 | 0.351
55 | GENSCREEN ULTRA HIV Ag-Ab | 8K0770 | 0.356 | 0.372 | 28.662 | 0.347 | 0.347 | 10.98 | 0.318 | 0.375
68 | GENSCREEN ULTRA HIV Ag-Ab | 9A0772 | 0.2554 | 0.2595 | 14.18 | 0.2108 | 0.2554 | 14.18 | 0.2554 | 0.223
112 | GENSCREEN ULTRA HIV Ag-Ab | 8H0123 | 0.1515 | 0.1645 | 12.645 | 0.1558 | 0.1688 | 12.5541 | 0.1602 | 0.1429
126 | GENSCREEN ULTRA HIV Ag-Ab |  | 0.229 | 0.212 | 11.98 | 0.2089 | 0.212 | 11.551 | 0.188 | 0.274
| Mean |  | 0.26338 | 0.2732 | 11.5346 | 0.21712 | 0.25904 | 11.4192 | 0.24732 | 0.30575
| Median |  | 0.2554 | 0.2595 | 11.98 | 0.20985 | 0.2554 | 11.2655 | 0.2554 | 0.3125
| S.Dev |  | 0.080837 | 0.090418 | 2.0906 | 0.056643 | 0.072303 | 1.7664 | 0.072025 | 0.069997
| C.V |  | 30.692 | 33.0957 | 18.1242 | 26.0879 | 27.912 | 15.4689 | 29.1221 | 22.8935
| Upper Limit |  | 0.42505 | 0.45404 | 15.7157 | 0.33041 | 0.40365 | 14.952 | 0.39137 | 0.44574
| Lower Limit |  | 0.10171 | 0.092365 | 7.3535 | 0.10384 | 0.11443 | 7.8863 | 0.10327 | 0.16576
97 | HIV TRDOT+ Ag. 4th GEN J.MITRA | HTA010119 | NONREACTIVE | NONREACTIVE | REACTIVE | NONREACTIVE | NONREACTIVE | REACTIVE | NONREACTIVE | NONREACTIVE
15 | HIV TRIDOT DIAG.ENT. | HTD091829 | HIV19/3: REACTIVE; HIV19/6: REACTIVE
58 | HIV TRIDOT DIAG.ENT. | HTD091828 | HIV19/3: REACTIVE; HIV19/6: REACTIVE
60 | HIV TRIDOT DIAG.ENT. | HTD041911 | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE
68 | HIV TRIDOT DIAG.ENT. | HTD111833 | HIV19/3: REACTIVE; HIV19/6: REACTIVE
69 | HIV TRIDOT DIAG.ENT. | HTD121835 | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE
96 | HIV TRIDOT DIAG.ENT. | HTD031907 | HIV19/3: REACTIVE; HIV19/6: REACTIVE
98 | HIV TRIDOT DIAG.ENT. | HTD030319 | HIV19/3: REACTIVE; HIV19/6: REACTIVE
52 | HIV TRIDOT DIAG.ENT. | HTD061922 | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE
55 | HIV TRIDOT DIAG.ENT. | HTD121835 | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE
74 | HIV TRIDOT DIAG.ENT. | HTD121836 | NEGATIVE | NEGATIVE | RECTIVE | NEGATIVE | NEGATIVE | RECTIVE | NEGATIVE | NEGATIVE
117 | HIV TRIDOT DIAG.ENT. | HTD041911 | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE | RECTIVE | NONREACTIVE | NONREACTIVE
15 | LABSYSTEM DIAG. TRIVITRON | HIVFE007 | 0.1172 | 0.0763 | 10.055 | 0.4127 | 0.0872 | 10.2727 | 0.0736 | 0.089
2 | MERILISA HIV 1-2 Gen 3 | MI201938 | 0.095 | 0.063 | 10.059 | 0.059 | 0.086 | 10.095 | 0.059 | 0.063
4 | MERILISA HIV 1-2 Gen 3 | MI 101825 | 0.146 | 0.176 | 9.708 | 0.15 | 0.206 | 10.459 | 0.107 | 0.163
9 | MERILISA HIV 1-2 Gen 3 | MI 0 11928 | 0.124 | 0.115 | 9.597 | 0.097 | 0.102 | 9.235 | 0.133 | 0.164
10 | MERILISA HIV 1-2 Gen 3 | MI1011928 | 0.046 | 0.046 | 17.545 | 0.064 | 0.06 | 17.358 | 0.064 | 0.069
26 | MERILISA HIV 1-2 Gen 3 | MI011928 | 0.085 | 0.098 | 7.156 | 0.085 | 0.13 | 7.941 | 0.094 | 0.098
30 | MERILISA HIV 1-2 Gen 3 | MI011928 | 0.09 | 0.157 | 8.649 | 0.209 | 0.177 | 7.952 | 0.169 | 0.165
33 | MERILISA HIV 1-2 Gen 3 | MI101864 | 0.113 | 0.097 | 9.105 | 0.097 | 0.168 | 9.164 | 0.118 | 0.13
34 | MERILISA HIV 1-2 Gen 3 | MI011938 | 0.26 | 0.24 | 12.77 | 0.17 | 0.37 | 11.98 | 0.3 | 0.19
41 | MERILISA HIV 1-2 Gen 3 | MI101850 | 0.099 | 0.163 | 8.801 | 0.129 | 0.185 | 9.672 | 0.099 | 0.051
41 | MERILISA HIV 1-2 Gen 3 | MI101850 | 0.168 | 0.141 | 9.566 | 0.11 | 0.159 | 9.898 | 0.119 | 0.101
45 | MERILISA HIV 1-2 Gen 3 | MI101827 | 0.08 | 0.09 | 13.31 | 0.12 | 0.11 | 13.24 | 0.13 | 0.09
52 | MERILISA HIV 1-2 Gen 3 | MI011928 | 0.132 | 0.0334 | 3.93 | 0.0271 | 0.0186 | 3.83 | 0.0728 | 0.29
64 | MERILISA HIV 1-2 Gen 3 | MI101826 | 0.2132 | 0.2182 | 16.662 | 0.2083 | 0.238 | 15.912 | 0.2331 | 0.2132
73 | MERILISA HIV 1-2 Gen 3 | MI011927 | 0.005 | 0.005 | 12.7 | 0.03 | 0.005 | 13.02 | 0.005 | 0.003
79 | MERILISA HIV 1-2 Gen 3 | MI021923 | 0.058 | 0.102 | 9.457 | 0.071 | 0.043 | 7.793 | 0.062 | 0.084
82 | MERILISA HIV 1-2 Gen 3 | MI101864 | 0.101 | 0.115 | 10.74 | 0.101 | 0.119 | 11.163 | 0.123 | 0.097
89 | MERILISA HIV 1-2 Gen 3 | MI041924 | 0.05213 | 0.04739 | 10.1706 | 0.03791 | 0.04739 | 10.8625 | 0.042654 | 0.0284
93 | MERILISA HIV 1-2 Gen 3 | MI011928 | 0.155 | 0.189 | 13.411 | 0.193 | 0.487 | 12.924 | 0.638 | 0.218
96 | MERILISA HIV 1-2 Gen 3 | MI011937 | 0.043 | 0.038 | 15.251 | 0.038 | 0.043 | 14.754 | 0.014 | 0.038
97 | MERILISA HIV 1-2 Gen 3 | MI041924 | 0.132 | 0.123 | 6.229 | 0.115 | 0.123 | 7.379 | 0.119 | 0.123
102 | MERILISA HIV 1-2 Gen 3 | MI101850 | 0.017 | 0.017 | 11.674 | 0.022 | 0.035 | 11.857 | 0.013 | 0.075
50 | MERILISA HIV 1-2 GEN 3 | MI011928 | 0.105 | 0.155 | 13.61 | 0.037 | 0.046 | 12.59 | 0.088 | 0.214
| Mean |  | 0.093363 | 0.094719 | 10.9313 | 0.075807 | 0.091199 | 10.5581 | 0.095364 | 0.084917
| Median |  | 0.099 | 0.1085 | 10.1148 | 0.097 | 0.1145 | 10.6608 | 0.103 | 0.0995
| S.Dev |  | 0.034295 | 0.041374 | 2.6723 | 0.035511 | 0.044891 | 2.7858 | 0.02918 | 0.027418
| C.V |  | 36.7329 | 43.6804 | 24.4463 | 46.8438 | 49.2232 | 26.3857 | 30.5991 | 32.2875
| Upper Limit |  | 0.16195 | 0.17747 | 16.2759 | 0.14683 | 0.18098 | 16.1298 | 0.15372 | 0.13975
| Lower Limit |  | 0.024773 | 0.011972 | 5.5867 | 0.0047853 | 0.0014169 | 4.9865 | 0.037003 | 0.030082
37 | MERILISA HIV Ag.& Ab. 4th GEN | MI011935 | 0.06 | 0.07 | 12.4 | 0.08 | 0.13 | 13.91 | 0.07 | 0.07
58 | MERILISA HIV Ag.& Ab. 4th GEN | MI051952 | 0.045 | 0.068 | 6.137 | 0.114 | 0.064 | 4.646 | 0.022 | 0.059
61 | MERILISA HIV Ag.& Ab. 4th GEN | MI051952 | 0.063 | 0.058 | 9.603 | 0.054 | 0.049 | 9.337 | 0.045 | 0.049
89 | MERSCREEN HIV 1-2 WB MERIL | MI081855 | HIV19/3: REACTIVE; HIV19/6: REACTIVE
40 | MICROLSA J.MITRA | EIV010319 | 0.036 | 0.0053 | 4.8794 | 0.0197 | 0.02 | 4 | 0.0071 | 0.0089
16 | MICROLSA J.MITRA Ag.& Ab. | EIA 020519 | 0.553 | 0.399 | 12.99 | 0.667 | 0.264 | 12.78 | 0.256 | 0.276
57 | MICROLSA J.MITRA Ag.& Ab. | EIA 020519 | 0.019 | 0.048 | 2.729 | 0.049 | 0.038 | 2.916 | 0.032 | 0.043
74 | MICROLSA J.MITRA Ag.& Ab. | EIA040818 | 0.12 | 0.06 | 8.9 | 0.16 | 0.041 | 7.87 | 0.06 | 0.07
122 | MICROLSA J.MITRA Ag.& Ab. | EIA 020519 | 0.328 | 0.306 | 9.91 | 0.164 | 0.213 | 11.293 | 0.24 | 0.182
74 | ORTHO VITROS ECI | 390 | 0.11 | 0.09 | 188 | 0.1 | 0.05 | 151 | 0.08 | 0.1
75 | ORTHO VITROS ECI | 390 | 0.1 | 0.07 | 166 | 0.08 | 0.08 | 169 | 0.07 | 0.1
98 | ORTHO VITROS ECI | 370 | 0.2 | 0.18 | 137 | 0.17 | 0.13 | 83.1 | 0.18 | 0.19
101 | ORTHO VITROS ECI | 390 | 0.25 | 0.24 | 173 | 0.23 | 0.19 | 170 | 0.21 | 0.23
104 | ORTHO VITROS ECI | 370 | 0.08 | 0.07 | 152 | 0.06 | 0.06 | 149 | 0.06 | 0.08
110 | ORTHO VITROS ECI | 370 | 0.18 | 0.08 | 146 | 0.1 | 0.08 | 136 | 0.08 | 0.12
120 | ORTHO VITROS ECI | 390 | 0.11 | 0.08 | 161 | 0.09 | 0.09 | 175 | 0.11 | 0.13
123 | ORTHO VITROS ECI | 360 | 0.11 | 0.13 | 200 | 0.1 | 0.08 | 201 | 0.08 | 0.1
11 | ORTHO VITROS ECI | 390 | 0.27 | 0.15 | 168 | 0.12 | 0.15 | 160 | 0.14 | 0.13
90 | ORTHO VITROS ECI | 370 | 0.14 |  | 132 | 0.15 | 0.12 | 136 | 0.12 | 0.15
91 | ORTHO VITROS ECI | 360 | 0.11 | 0.08 | 162 | 0.08 | 0.13 | 179 | 0.14 | 0.16
121 | ORTHO VITROS ECI | 360 | 0.15 | 0.11 | 179 | 0.11 | 0.13 | 176 | 0.13 | 0.14
128 | ORTHO VITROS ECI | 360 | 0.16 | 0.08 | 194 | 0.11 | 0.12 | 198 | 0.18 | 0.11
113 | ORTHO VITROS ECI | 360 | 0.16 | 0.1 | 171 | 0.13 | 0.14 | 173 | 0.17 | 0.25
| Mean |  | 0.165 | 0.094545 | 166.3571 | 0.12111 | 0.12333 | 167.1538 | 0.125 | 0.12
| Median |  | 0.15 | 0.105 | 171 | 0.11 | 0.13 | 176 | 0.14 | 0.14
| S.Dev |  | 0.054681 | 0.025832 | 20.1975 | 0.024721 | 0.024221 | 20.424 | 0.047515 | 0.024495
| C.V |  | 33.1399 | 27.3219 | 12.1411 | 20.4116 | 19.6388 | 12.2187 | 38.0121 | 20.4124
| Upper Limit |  | 0.27436 | 0.14621 | 206.7522 | 0.17055 | 0.17178 | 208.0019 | 0.22003 | 0.16899
| Lower Limit |  | 0.055638 | 0.042882 | 125.9621 | 0.07167 | 0.074891 | 126.3058 | 0.02997 | 0.07101
117 | OSCAR MEDICARE | HIV4EL-010 | 0.2124 | 0.1983 | 11.6465 | 0.2232 | 0.1871 | 12.8112 | 0.1626 | 0.1654
85 | QUALISA QULPRO TULIP | 49188 | 0.208 | 0.323 | 10.169 | 0.226 | 0.329 | 10.234 | 0.303 | 0.216
87 | QUALISA QULPRO TULIP | 49180 | 0.011 | 0.255 | 12.962 | 0.04 | 0.003 | 12.962 | 0.355 | 0.351
16 | QUALISA QULPRO TULIP | 54393 | 0.033 | 0.004 | 13.112 | 0.004 | 0.0625 | 14.99 | 0.0384 | 0.067
36 | QUALISA QULPRO TULIP | 49188 | 0.11 | 0.16 | 10.87 | 0.2 | 0.23 | 10.33 | 0.05 | 0.02
56 | QUALISA QULPRO TULIP | 49187 | 0.141 | 0.114 | 13.882 | 0.136 | 0.123 | 13.923 | 0.136 | 0.105
60 | QUALISA QULPRO TULIP | 49189 | 0.141 | 0.22 | 13.137 | 0.163 | 0.15 | 12.115 | 0.009 | 0.022
69 | QUALISA QULPRO TULIP | 49186 | 0.101 | 0.106 | 20.85 | 0.131 | 0.035 | 20.6 | 0.23 | 0.091
129 | QUALISA QULPRO TULIP | 49191 | 0.07 | 0.07 | 8.744 | 0.07 | 0.065 | 8.107 | 0.065 | 0.07
| Mean |  | 0.099333 | 0.15391 | 11.8394 | 0.108 | 0.0871 | 11.8087 | 0.07235 | 0.0625
| Median |  | 0.101 | 0.114 | 13.112 | 0.131 | 0.065 | 12.962 | 0.065 | 0.07
| S.Dev |  | 0.04207 | 0.071488 | 1.9169 | 0.051005 | 0.047574 | 2.3939 | 0.043808 | 0.035047
| C.V |  | 42.3521 | 46.4483 | 16.1909 | 47.2268 | 54.6202 | 20.2721 | 60.5504 | 56.0754
| Upper Limit |  | 0.18347 | 0.29689 | 15.6732 | 0.21001 | 0.18225 | 16.5965 | 0.15997 | 0.13259
| Lower Limit |  | 0.015194 | 0.010933 | 8.0056 | 0.0059902 | -0.0080483 | 7.021 | -0.015266 | -0.0075942
87 | QUALPRO DIAG. | 48426 | HIV19/3: REACTIVE; HIV19/6: REACTIVE
64 | RETROQUIC TULIP DIAG. | 541241 | N.R | N.R | REACTIVE | N.R | N.R | REACTIVE | N.R | N.R
18 | ROCHE COBAS E411 | 391001 | 0.327 | 0.262 | 772.11 | 0.27 | 0.27 | 760.33 | 0.281 | 0.293
18 | ROCHE COBAS E411 | 400139 | 0.295 | 0.236 | 694.9 | 0.243 | 0.243 | 684.3 | 0.253 | 0.264
43 | ROCHE COBAS E411 | 391001 | 0.297 | 0.325 | 776.2 | 0.298 | 0.295 | 711.9 | 0.283 | 0.669
95 | ROCHE COBAS E411 | 391001 | 0.318 | 0.294 | 567.7 | 0.311 | 0.302 | 636.5 | 0.29 | 0.304
95 | ROCHE COBAS E411 | 391001 | HIV19/3: 559.1; HIV19/6: 625.4
| Mean |  | 0.30925 | 0.27925 | 674.002 | 0.2805 | 0.2775 | 683.686 | 0.27675 | 0.3825
| Median |  | 0.3075 | 0.278 | 694.9 | 0.284 | 0.2825 | 684.3 | 0.282 | 0.2985
| S.Dev |  | 0.015756 | 0.038638 | 106.0765 | 0.030293 | 0.026789 | 55.4403 | 0.016297 | 0.19174
| C.V |  | 5.0949 | 13.8364 | 15.7383 | 10.7996 | 9.6538 | 8.109 | 5.8886 | 50.1291
| Upper Limit |  | 0.34076 | 0.35653 | 886.1551 | 0.34109 | 0.33108 | 794.5665 | 0.30934 | 0.76599
| Lower Limit |  | 0.27774 | 0.20197 | 461.8489 | 0.21991 | 0.22392 | 572.8055 | 0.24416 | -0.00098751
40 | SIGNAL ARKRAY DIAG. | 4000020643 | NAGATIVE | NAGATIVE | POSITIVE | NAGATIVE | NAGATIVE | POSITIVE | NAGATIVE | NAGATIVE
61 | TRIDOT TULIP DIAG. | HT0051915 | NEGATIVE | NEGATIVE | POSITIVE | NEGATIVE | NEGATIVE | POSITIVE | NEGATIVE | NEGATIVE